## Report CAL.
[12 empty] | Certificate No. |  |  |  | BSN0001
[18 empty] | DATA RECORD FORM OF |  |  |  | INFUSION PUMP
Certificate of Calibration
[18 empty] | Site | WARD 9/1 | #N/A
| EQUIPMENT |  |  |  |  | : | INFUSION PUMP | [10 empty] | No. | BSN0001
| ID CODE |  |  |  |  | : | BSN0090 | [15 empty] | 2014
| MANUFACTURER |  |  |  |  | : | TOP | [10 empty] | EQUIPMENT ID | BSN0090 |  | PERFORM DATE | 2DEC14 |  | MASTER | Cal. |  | Pm. |  | DATA OF |  | RESULT | [18 empty] | No. | ADD | ID CODE | Cer.No | PM.No | MEDICAL EQUIPMENT | BRAND | MODEL | SERIAL NUMBER | DEPARTMENT | MAJOR | ACTION | DATE | ENGINEER | STATUS | REMARK
| MODEL |  |  |  |  | : | TOP-2200 | [10 empty] | EQUIPMENT | INFUSION PUMP |  | DUE DATE | 2015-12-01 |  | 1 | - |  |  |  | PM&CAL |  | OUT OF RANGE | [20 empty] | BSN0090 |  |  | INFUSION PUMP | TOP | TOP-2200 | EL17030S | WARD 9/1 |  | PM&CAL | 2DEC14 | 518096 | OUT OF RANGE | 0
| SERIAL No. |  |  |  |  | : | EL17030S | [10 empty] | MANUFACTURER | TOP |  | ENGINEER | 518096 |  | 2
| LOCATION |  |  |  |  | : | WARD 9/1 | [10 empty] | MODEL | TOP-2200 |  | ENVIRONMENT | 24 | 0 C | 3 |  |  |  |  | PM&CAL |  | PASS
| SUBMITTED BY |  |  |  |  | : | #N/A | [10 empty] | S/N | EL17030S |  |  | 60 | %RH | 4 |  |  |  |  | CAL |  | OUT OF RANGE
[7 empty] | #N/A | [10 empty] | DEPT./HOSPITAL | WARD 9/1 |  |  |  |  | 5 |  |  |  |  | PM |  | DEFECTIVE
[7 empty] | #N/A | [16 empty] | 6
| CALIBRATED DATE |  |  |  |  | : | 2DEC14
| ISSUE DATE |  |  |  |  | : | 2014-12-07
| Performed by : | [8 empty] | Approved by :
|  |  | Panot | [8 empty] | Kumpee
This certificate may not be reproduced except in full unless permission for reproduction has been
obtained in writing from the calibration.
CONDITION OF THIS RESULT OF TEST
1. REFERENCE STANDARD INSTRUMENT :
MASTER |  |  |  | MANUFACTURER |  |  |  | MODEL |  |  | SERIAL NO. |  |  | CERTIFICATE NO. |  |  | DUE DATE
- |  |  |  | - |  |  |  | - |  |  | - |  |  | - |  |  | -
2. THIS CERTIFICATION IS TRACEABLE TO :
-
3. THIS RESULT OF TEST WAS FOUND ACCURATE AS SHOW ON DATE AND PLACE OF
TEST ONLY
[12 empty] | Certificate No. |  |  |  | BSN0001
Calibration Report
| EQUIPMENT | [5 empty] | : | INFUSION PUMP
| ID CODE | [5 empty] | : | BSN0090
| MANUFACTURER | [5 empty] | : | TOP
| MODEL | [5 empty] | : | TOP-2200
| SERIAL No. | [5 empty] | : | EL17030S
| DATE OF CALIBRATION | [5 empty] | : | 2DEC14
| ENVIRONMENT | [5 empty] | : | TEMPERATURE |  |  |  |  | 24 | °C
[8 empty] | RELATIVE HUMIDITY |  |  |  |  | 60 | %
PROCEDURE USED :
This instrument was calibration by comparison with standard
| MEASURMENT RESULT : | [30 empty] | UNCERTAINTY :2014 |  |  |  | Last Update: | 2014-07-15
[32 empty] | INPT001 |  |  |  | used to | 2015-07-14
|  |  | X | Without Adjustment |  |  |  |  | Before Adjustment |  |  |  |  | After Adjustment
|  | Flow Rate | [29 empty] | Flow Rate
|  | UUC* Setting |  |  | Standard Reading |  |  |  | Error |  |  | % Error |  |  | Uncertainty |  |  | Parameter | Flow Rate | Unit | ml/hr | Tolerance |  |  |  | ± | 5 | % | Average | U. | Result(Include U.) | u1 | u2 | u3 | u4 | u5 | u6 | u7 | U1 | U2 | U3 | U4 | U5 | U6 | U7 | Uc | U. | veff. | K
|  | ml/hr |  |  | ml/hr |  |  |  | ml/hr |  |  | (%) |  |  | ml/hr |  |  | Setting | Reading 1 | Reading 2 | Reading 3 | Reading 4 | Min. |  |  | Max.
|  | 5 |  |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0! |  |  | #DIV/0! | #DIV/0! |  | 5 |  |  |  |  | 4.75 |  |  | 5.25 |  |  | #DIV/0! | #DIV/0! | #DIV/0! |  | 0.01 | 0.01 | 0.5 |  |  |  | #DIV/0! | 0.0028868 | 0.0028868 | 0.014434 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 80 |  |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0! |  |  | #DIV/0! | #DIV/0! |  | 80 |  |  |  |  | 76 |  |  | 84 |  |  | #DIV/0! | #DIV/0! | #DIV/0! |  | 0.01 | 0.01 | 0.5 |  |  |  | #DIV/0! | 0.0028868 | 0.0028868 | 0.23094 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 200 |  |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0! |  |  | #DIV/0! | #DIV/0! |  | 200 |  |  |  |  | 190 |  |  | 210 |  |  | #DIV/0! | #DIV/0! | #DIV/0! |  | 0.01 | 0.01 | 0.5 |  |  |  | #DIV/0! | 0.0028868 | 0.0028868 | 0.57735 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | Occlusion Pressure | [29 empty] | Occlusion Pressure
|  | UUC* Setting |  |  | Standard Reading | [12 empty] | Parameter | Occlusion Pressure | Unit | mmHg | Tolerance | < 1,200 mmHg | [5 empty] | Average | U. | Result(Include U.) | u1 | u2 | u3 | u4 | u5 | u6 | u7 | U1 | U2 | U3 | U4 | U5 | U6 | U7 | Uc | U. | veff. | K
|  | mmHg |  |  | mmHg | [12 empty] | Setting | Reading 1 | Reading 2 | Reading 3 | Reading 4 | Min. |  |  | Max.
|  | 1200 |  |  | #DIV/0! | [12 empty] | 1200 |  |  |  |  | N/A |  |  | 1200 |  |  | #DIV/0! | #DIV/0! | #DIV/0! |  | 1 | 0.1 |  | 1.2 |  |  | #DIV/0! | 0.28868 | 0.028868 | 0 | 0.6 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | Use IV Set Brand |  |  |  | 0 | [5 empty] | Drops/ml |  |  | 0 |  |  | Use IV Set Brand |  | Drops/ml
Comment
UUC* : Unit Under Calibration
The reported uncertainty of measurement was based on a standard uncertainty multiplied
by a coverage factor k = 2, providing a level of confidence of approximately 95%
[43 empty] | U1 = Uncertainty Type A |  |  |  |  | t-distribution Table
[18 empty] | FI-BME-NHS-CP-001/2 แก้ไขครั้งที่ 00 วันที่บังคับใช้ 01/09/2557 | [24 empty] | U2 = Resolution of UUC | [5 empty] | k
[43 empty] | U3 = Resolution of STD |  |  |  |  | 1 | 13.97
[43 empty] | U4 = Accuracy ( of reading ) |  |  |  |  | 2 | 4.53
[43 empty] | U5 = Uncertainty of STD |  |  |  |  | 3 | 3.31
[43 empty] | U6 = Accuracy ( +/- Value ) |  |  |  |  | 4 | 2.87
[43 empty] | U7=Driff of standard |  |  |  |  | 5 | 2.65
[48 empty] | 6 | 2.52
[48 empty] | 7 | 2.43
[48 empty] | 8 | 2.37
[48 empty] | 9 | 2.32
[48 empty] | 10 | 2.28
[48 empty] | 11 | 2.25
[48 empty] | 12 | 2.23
[48 empty] | 13 | 2.21
[48 empty] | 14 | 2.2
[48 empty] | 15 | 2.18
[48 empty] | 16 | 2.17
[48 empty] | 17 | 2.16
[48 empty] | 18 | 2.15
[48 empty] | 19 | 2.14
[48 empty] | 20 | 2.13
[48 empty] | 25 | 2.11
[48 empty] | 30 | 2.09
[48 empty] | 35 | 2.07
[48 empty] | 40 | 2.06
[48 empty] | 45 | 2.06
[48 empty] | 50 | 2.05
[48 empty] | 100 | 2.025
[48 empty] | ∞ | 2
Ref : M3003 Page 29
Note on AP68: Uncertainty Type A
Note on AQ68: Resolution of UUC
Note on AR68: Resolution of STD
Note on AS68: Accuracy ( of reading )
Note on AT68: Uncertainty of STD
Note on AU68: Accuracy ( +/- Value )
Note on AV68: Drift of standard
Note on AP74: Uncertainty Type A
Note on AQ74: Resolution of UUC
Note on AR74: Resolution of STD
Note on AS74: Accuracy ( of reading )
Note on AT74: Uncertainty of STD
Note on AU74: Accuracy ( +/- Value )
Note on AV74: Drift of standard